## Notes
Goals
Create evidence-based models for a set of patient diagnoses
Validate output from Pulse Kitware Physiology Simulator using published data
A. Select references that give values for ventilator settings and resultant gas exchange variables
B. Empirically select ventilator settings within reference ranges to get physiology model to be in reference ranges
C. Modify physiology model and repeat step B as needed to get outputs into reference ranges
Patient Diagnoses
Normal
ARDS (mild, moderate, severe)
COPD (mild, moderate, severe)
Modes of Ventilation
PC-IMV
VC-CMV
Procedure for Validating Ventilation
Establish Lo and Hi values for PaCO2 from reference #1 (range is better than SD because normal distribution not assumed)
For normal and COPD use vent settings from reference #1
For ARDS, set test values for tidal volume and rate from reference #5 (has wider range of minute ventilation than Arnal so wider range of dead space of patients was included in dataset)
Run physiology simulation
If PaCO2 is out of limits then adjust rate using equation from reference #6
new rate = (old PaCO2 * old rate)/(AVERAGE(low limit PaCO2,high limit PaCO2))
new tidal volume = (old PaCO2 * old tidal volume)/(AVERAGE(low limit PaCO2,high limit PaCO2))
Run physiology simulation and confirm new PaCO2 is within limits
Procedure for Validating Oxygenation
Set test values for FiO2, tidal volume and rate from - Arnal J-M, Garnero A, Novonti D, et al. Feasibility study on full closed-loop control ventilation (IntelliVent-ASVTM) in ICU patients with acute respiratory failure: a prospective observational comparative study. Crit Care. 2013;17(5):R196.
Set limits for PaO2 (ARDS patients) from reference #4
-calculate limits: PaO2 = (Oxygenation Ratio +/- 1 SD) x FiO2
Run physiology simulation
If PaO2 is out of limits then adjust FiO2 using equation from reference #6: new FiO2 = AVERAGE(Low PaO2,Hi PaO2)/(old PaO2/old FiO2)
Run physiology simulation and confirm new PaO2 is within limits
References
1. Arnal J-M, Garnero A, Novonti D, et al. Feasibility study on full closed-loop control ventilation (IntelliVent-ASVTM) in ICU patients with acute respiratory failure: a prospective observational comparative study. Crit Care. 2013;17(5):R196.
2. Arnal JM, Garnero A, Saoli M, Chatburn RL. Parameters for Simulation of Adult Patients During Mechanical Ventilation. Respir Care. 2017;63(2):158-168.
3. Ferguson, N. D., Fan, E., Camporota, L., Antonelli, M., Anzueto, A., Beale, R., ... & Rhodes, A. (2012). The Berlin definition of ARDS: an expanded rationale, justification, and supplementary material. Intensive care medicine, 38(10), 1573-1582.
4. El-Khatib MF, Bouakl IJ, Ayoub CM, et al. Comparison of the Oxygenation Factor and the Oxygenation Ratio in Subjects With ARDS. Respir Care. July 2020.
5. Ventilation with lower tidal volumes as compared with traditional tidal volumes for acute lung injury and the acute respiratory distress syndrome. The Acute Respiratory Distress Syndrome Network. N Engl J Med. 2000;342(18):1301-1308.
6. Kacmarek RM, Stoller JK, Heuer AJ. Egan's fundmentals of respiratory care. 12th edition. St. Louis: Elsevier, 2021:1091-1097.

## Model
ARDS & COPD Validation
Patient1
| Sex | male |  | Goal
| Age (years) | 44 |  | Validate performance of physiology simulator for passive and active patient models in two modes of ventilation -Note that expected values are for PC-CMV, but values for VC-CMV should be within specified ranges -Expected values for SpO2 are calculated -Active effort = spontaneous breathing; passive effort = apnea
| Predicted Weight (kg) | 75.3
| RR Baseline (bpm) | 12
| Pathology | Normal
| Interventions2 | Invasive Mechanical Ventilation
NORMAL
|  |  |  | Simulation Parameters1 |  |  | Expected Values2
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
| VT Target (mL/kg) | 7 |  | passive | normal | 0 | 15 | 21 | 0.4 | 0.8 | 12 | 17 | 20 | 54 | 7.27 | 7.4 | 35 | 42 | 85 | 11.3324 | 114 | 15.1987 | 264 | 366 | 96.4005 | 98.4621
| VT Target (mL) | 527.1 |  |  |  | Observed Values from Physiology Simulator | 15 |  | 0.6 |  | 13 |  | 50 |  | 7.4 |  | 40 |  | 104 |  |  |  | 344 |  | 98.2
| PEEP (cm H2O) | 5 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| FiO2 (%) | 30
| Mandatory Rate (bpm) | 15
| I:E | 0.6 |  |  |  |  | Comments
| Minute Ventilation (L/min) | 7.9065
| Mode | VC-CMV |  | Simulation Parameters1 |  |  | Expected Values2
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| VT Target (mL/kg) | 7 |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
| VT Target (mL) | 527.1 |  | passive | normal | 0 | 15 | 20 | 0.4 | 0.8 | 12 | 17 | 20 | 54 | 7.27 | 7.4 | 35 | 42 | 85 | 11.3324 | 114 | 15.1987 | 264 | 366 | 96.4005 | 98.4621
| PEEP (cm H2O) | 5 |  |  |  | Observed Values from Physiology Simulator | 15 |  | 0.6 |  | 13 |  | 50 |  | 7.4 |  | 40 |  | 104 |  |  |  | 344 |  | 98.2
| FiO2 (%) | 30 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| Mandatory Rate (bpm) | 15
| I:E | 0.6 |  |  |  |  | Commnets
| Minute Ventilation (L/min) | 7.9065
| Mode | VC-AC |  | Simulation Parameters1 |  |  | Expected Values2
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| VT Target (mL/kg) | 7 |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
| VT Target (mL) | 527.1 |  | active | normal | 0 | 15 | 20 | 0.4 | 0.8 | 12 | 17 | 20 | 54 | 7.27 | 7.4 | 35 | 42 | 85 | 11.3324 | 114 | 15.1987 | 264 | 366 | 96.4005 | 98.4621
| PEEP (cm H2O) | 5 |  |  |  | Observed Values from Physiology Simulator | 15 |  | 0.6 |  | 13 |  | 50 |  | 7.4 |  | 40 |  | 99 |  |  |  | 331 |  | 98
| FiO2 (%) | 30 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| Mandatory Rate (bpm) | 0
| I:E | 0.6 |  |  |  |  | Commnets
| Minute Ventilation (L/min) | 7.9065
ARDS
|  |  |  | Simulation Parameters1 |  |  | Expected Values2,3
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
| VT Target (mL/kg) | 6 |  | passive | ARDS | 0.3 | 15 | 35 | 0.8 | 1 | 10 | 16 | 20 | 54 | 7.15 | 7.34 | 45 | 57 | 103.4 | 13.7855 | 132.54 | 17.6705 | 200 | 300 | 97.9547 | 99.013
| VT Target (mL) | 451.8 |  |  |  | Observed Values from Physiology Simulator | 16.8627 |  | 0.9 |  | 13 |  | 40 |  | 7.3 |  | 50 |  | 107 |  |  |  | 265 |  | 98
| PEEP (cm H2O) | 11 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| FiO2 (%) | 40 | range from Ref #4: 47 ± 14
| Mandatory Rate (bpm) | 16.8627 |  |  |  |  | Commnets | [17 empty] | target
| I:E | 0.9 |  |  |  |  | If PaO2 is out of limits then adjust FiO2 using equation from reference #6: new FiO2 = AVERAGE(Low PaO2,Hi PaO2)/(old PaO2/old FiO2) | [13 empty] | New FiO2 | PCO2 | Old FiO2 | Old PaO2 | PaO2
| Minute Ventilation (L/min) | 7.6186 | [18 empty] | 40.6897 | 55 | 30 | 87 | 118
|  |  |  | Simulation Parameters1 |  |  | Expected Values2,3
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
| VT Target (mL/kg) | 6 |  | passive | ARDS | 0.6 | 15 | 35 | 0.8 | 1 | 10 | 16 | 20 | 54 | 7.15 | 7.34 | 45 | 57 | 81.18 | 10.8231 | 116.82 | 15.5747 | 100 | 200 | 95.8976 | 98.5685
| VT Target (mL) | 451.8 |  |  |  | Observed Values from Physiology Simulator | 20 |  | 0.9 |  | 13 |  | 30 |  | 7.3 |  | 50 |  | 95 |  |  |  | 135 |  | 97.1
| PEEP (cm H2O) | 11 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| FiO2 (%) | 70
| Mandatory Rate (bpm) | 20
| I:E | 0.9 |  |  |  |  | Commnets
| Minute Ventilation (L/min) | 9.036
|  |  |  | Simulation Parameters1 |  |  | Expected Values2,3
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
| VT Target (mL/kg) | 6 |  | passive | ARDS | 0.9 | 15 | 35 | 0.8 | 1 | 10 | 16 | 20 | 54 | 7.15 | 7.34 | 45 | 57 | 56.07 | 7.4754 | 82.77 | 11.0351 | 0 | 100 | 88.7517 | 96.1174
| VT Target (mL) | 451.8 |  |  |  | Observed Values from Physiology Simulator | 25.4902 |  | 0.9 |  | 13 |  | 20 |  | 7.25 |  | 57 |  | 62 |  |  |  | 62 |  | 90.3
| PEEP (cm H2O) | 11 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| FiO2 (%) | 100
| Mandatory Rate (bpm) | 25.4902
| I:E | 0.9 |  |  |  |  | Commnets
| Minute Ventilation (L/min) | 11.5165
COPD
|  |  |  | Simulation Parameters1 |  |  | Expected Values2
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
Prior | Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
8.2 | VT Target (mL/kg) | 9 |  | passive | COPD | 0.3 | 12 | 15 | 0.32 | 0.5 | 17 | 26 | 46 | 61 | 7.2 | 7.31 | 44 | 67 | 77 | 10.2658 | 94 | 12.5323 | 206 | 295 | 95.2375 | 97.3036
| VT Target (mL) | 677.7 |  |  |  | Observed Values from Physiology Simulator | 12 |  | 0.43 |  | 17 |  | 57 |  | 7.27 |  | 58 |  | 93 |  |  |  | 295 |  | 96.5
| PEEP (cm H2O) | 8 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
| FiO2 (%) | 31 | [11 empty] | Prior | 7.25 |  | 60 |  | 90 |  |  |  | 290 |  | 96
13 | Mandatory Rate (bpm) | 11.8046 |  |  |  |  | Commnets | [9 empty] | target | [7 empty] | target | [6 empty] | target
| I:E | 0.43 |  |  |  |  | new VT is now in published range | [8 empty] | pH | PCO2 | HCO3 |  |  | New MV | PCO2 | Old MV | Old PaCO2 | PaCO2 |  |  | new PaCO2 | PCO2 | old VT | Old PaCO2 | new VT
8 | Minute Ventilation (L/min) | 8 | range from Ref#1: 5.3 - 9.0 | [12 empty] | 7.2726 | 56 | 25 |  |  | 8.6486 | 55 | 8 | 60 | 55.5 |  |  | 54.6667 | 55 | 7.2 | 60 | 96.2706
|  |  |  | Simulation Parameters1 |  |  | Expected Values2
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
8 | VT Target (mL/kg) | 8.5 |  | passive | COPD | 0.6 | 12 | 15 | 0.32 | 0.5 | 17 | 26 | 46 | 61 | 7.2 | 7.31 | 44 | 67 | 77 | 10.2658 | 94 | 12.5323 | 206 | 295 | 95.2375 | 97.3036
| VT Target (mL) | 640.05 |  |  |  | Observed Values from Physiology Simulator | 14 |  | 0.43 |  | 21 |  | 59 |  | 7.27 |  | 59 |  | 83 |  |  |  | 267 |  | 95.2
| PEEP (cm H2O) | 8 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Fail
| FiO2 (%) | 31 | [11 empty] | Prior | 7.17 |  | 73 |  | 80 |  |  |  | 257 |  | 93
12 | Mandatory Rate (bpm) | 14 | [14 empty] | target | [7 empty] | target
| I:E | 0.43 |  |  |  |  | Commnets | [8 empty] | pH | PCO2 | HCO3 |  |  | New MV | PCO2 | Old MV | Old PaCO2 | PaCO2
7.2 | Minute Ventilation (L/min) | 8.9607 | range from Ref#1: 5.3 - 9.0 |  |  |  | new minute ventilation = (old PaCO2 * old mintute ventilation) / target PaCO2 | [8 empty] | 7.2726 | 56 | 25 |  |  | 9.4703 | 55 | 7.2 | 73 | 55.5
reference #1 gives incompatible ranges for VT, RR, and MV, so choose VT and MV as valid ranges
|  |  |  | Simulation Parameters1 |  |  | Expected Values2
| Mode | PC-CMV |  | PHYSIOLOGY MODEL |  |  | RR (bpm) |  | I:E |  | R (cmH2O-s/L) |  | Cstat (mL/cmH2O) |  | pH |  | PaCO2 (mmHg) |  | PaO2 (mmHg) |  |  |  | PaO2/FiO2 (mmHg) |  | SpO2 (%)
| Inspiratory Pressure Target | set empirically to achieve VT Target |  | Effort | Dx | Severity | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo | Hi | Lo |  | Hi |  | Lo | Hi | Lo | Hi
7 | VT Target (mL/kg) | 9 |  | passive | COPD | 0.9 | 12 | 15 | 0.32 | 0.5 | 17 | 26 | 46 | 61 | 7.2 | 7.31 | 44 | 67 | 77 | 10.2658 | 94 | 12.5323 | 206 | 295 | 95.2375 | 97.3036
| VT Target (mL) | 677.7 |  |  |  | Observed Values from Physiology Simulator | 15 |  | 0.43 |  | 26 |  | 61 |  | 7.31 |  | 51 |  | 84 |  |  |  | 208 |  | 95.9
| PEEP (cm H2O) | 8 |  |  |  | Outcome | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  | Pass |  |  |  | Pass |  | Pass
35 | FiO2 (%) | 40 | [11 empty] | Prior | 7.14 |  | 80 |  | 77 |  |  |  | 212 |  | 91
13 | Mandatory Rate (bpm) | 15 | [14 empty] | target | [7 empty] | target
| I:E | 0.43 |  |  |  |  | Commnets | [8 empty] | pH | PCO2 | HCO3 |  |  | New MV | PCO2 | Old MV | Old PaCO2 | PaCO2
6.6 | Minute Ventilation (L/min) | 10.1655 | range from Ref#1: 5.3 - 9.0 |  |  |  | new minute ventilation = (old PaCO2 * old mintute ventilation) / target PaCO2 | [8 empty] | 7.2726 | 56 | 25 |  |  | 9.5135 | 55 | 6.6 | 80 | 55.5
reference #1 gives incompatible ranges for VT, RR, and MV, so choose VT and MV as valid ranges
If SpO2 is out of limits then adjust FiO2 using equation from reference #6: new FiO2 = AVERAGE(Low PaO2,Hi PaO2)/(old PaO2/old FiO2)
References
| 1. Arnal J-M, Garnero A, Novonti D, et al. Feasibility study on full closed-loop control ventilation (IntelliVent-ASVTM) in ICU patients with acute respiratory failure: a prospective observational comparative study. Crit Care. 2013;17(5):R196. | [25 empty] | arnal2013feasibility
| 2. Arnal JM, Garnero A, Saoli M, Chatburn RL. Parameters for Simulation of Adult Patients During Mechanical Ventilation. Respir Care. 2017;63(2):158-168. | [25 empty] | arnal2018parameters
| 3. Ferguson, N. D., Fan, E., Camporota, L., Antonelli, M., Anzueto, A., Beale, R., ... & Rhodes, A. (2012). The Berlin definition of ARDS: an expanded rationale, justification, and supplementary material. Intensive care medicine, 38(10), 1573-1582. | [25 empty] | ferguson2012berlin
| 4. El-Khatib MF, Bouakl IJ, Ayoub CM, et al. Comparison of the Oxygenation Factor and the Oxygenation Ratio in Subjects With ARDS. Respir Care. July 2020. | [25 empty] | el2020comparison
| 5. Ventilation with lower tidal volumes as compared with traditional tidal volumes for acute lung injury and the acute respiratory distress syndrome. The Acute Respiratory Distress Syndrome Network. N Engl J Med. 2000;342(18):1301-1308. | [25 empty] | acute2000ventilation
| 6. Kacmarek RM, Stoller JK, Heuer AJ. Egan's fundmentals of respiratory care. 12th edition. St. Louis: Elsevier, 2021:1091-1097. | [25 empty] | kacmarek2016egan
Reference #2
Reference #1
Reference #4
Reference #5
Note on G16: diagnosis
Note on H16: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on G28: diagnosis
Note on H28: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on G39: diagnosis
Note on H39: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on G51: diagnosis
Note on H51: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on J52: in range from referemce #5
Note on G63: diagnosis
Note on H63: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on J64: in range from referemce #5
Note on G75: diagnosis
Note on H75: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on J76: in range from referemce #5
Note on D79: within upper limit of reference #5
Note on G89: diagnosis
Note on H89: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on G101: diagnosis
Note on H101: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9
Note on D102: arbitrary but within limits reference 1
Note on G113: diagnosis
Note on H113: normal = 0 mild = 0.3 moderate = 0.6 Severe = 0.9